## README
Novaflow RCA V4 — Sterilizer Triple Peak · MAE+RMSE Scoring Engine
Scoring pipeline: Full-cycle normalisation → point-by-point MAE & RMSE vs benchmark → combined_error = 0.5×MAE + 0.5×RMSE → score = 100 × e^(−2.2 × combined_error). RCA layer explains WHY MAE/RMSE is high in each stage sub-window.
Sheet | Purpose
1_SCORING_PIPELINE | Full formula reference. Normalisation → MAE/RMSE → combined_error → exponential score. k=2.2, weights 50/50. Score table showing error→score mapping. Sub-window decomposition method for RCA.
2_STAGE_DEFINITION | Triple peak window boundaries using pressure events (not fixed time). S1=first peak+release, S2=second peak+release, S3=third peak+hold. Each stage references its own normalised benchmark sub-array from InfluxDB.
3_RULES_MASTER | All RCA rules with correct MAE/RMSE-based formulas. Each rule reads a sub-window MAE (e.g. MAE_s1_ramp, MAE_s3_hold) and compares against threshold. Divergence pattern A–F. Attribution group P1–P5.
4_THRESHOLD_LIBRARY | All MAE/RMSE thresholds in one place. Thresholds are on 0–1 normalised scale. Edit here to tune sensitivity.
5_ATTRIBUTION_ENGINE | P1→P5 priority decision tree. Each step checks which sub-window MAE is elevated and how many sterilizers are affected. Programmer implements as ordered if-else.
6_SCORE_MODEL | Weights S1=20%, S2=30%, S3=50%. Group deduction caps. Score bands with oil loss context. Cycle score formula.
7_RECOMMENDATION_MATRIX | KEY OUTPUT: (rec_key) → recommendation text EN + BM. Programmer does one lookup after attribution. Includes benchmark gap context.
8_EXHAUST_SAFETY_CHECK | Pass/fail only. Does NOT affect score. Safety alarm if residual pressure detected.

## 1_SCORING_PIPELINE
Scoring Pipeline — Complete Formula Reference (k=2.2, weights 50/50)
STEP 1 — FULL CYCLE NORMALISATION
Normalise BOTH actual sensor data and benchmark across the full cycle (not per-stage). This ensures the same scale across factories regardless of absolute pressure values. P_norm(t) = [ P_actual(t) − min(P_full_cycle) ] / [ max(P_full_cycle) − min(P_full_cycle) ] B_norm(t) = [ B_benchmark(t) − min(B_full_cycle) ] / [ max(B_full_cycle) − min(B_full_cycle) ] Result: both P_norm and B_norm are on 0–1 scale. Ready for point-by-point comparison.
STEP 2 — POINT-BY-POINT ERROR (per stage sub-window)
After normalisation, slice into stage windows. Compute MAE and RMSE per stage. For each stage window [t_start → t_end]: n = number of time points in window error(t) = P_norm(t) − B_norm(t) ← signed difference at each point MAE_stage = (1/n) × Σ ¦error(t)¦ ← mean absolute error RMSE_stage = √[ (1/n) × Σ error(t)² ] ← root mean square error MAE treats all deviations equally. RMSE penalises large spikes more (squared term). Captures pressure oscillation better. Both on 0–1 scale (since P_norm and B_norm are 0–1).
STEP 3 — COMBINED ERROR
Combine MAE and RMSE with equal 50/50 weight: combined_error_stage = 0.5 × MAE_stage + 0.5 × RMSE_stage Range: 0 (perfect match) to 1 (maximum deviation). combined_error is the single value fed into the scoring function.
STEP 4 — EXPONENTIAL SCORE CONVERSION (k = 2.2)
Convert combined_error to 0–100 score using exponential decay: stage_score = 100 × e^(−2.2 × combined_error_stage) Score table for reference (k=2.2): combined_error = 0.00 → score = 100.0 (perfect match) combined_error = 0.05 → score = 89.6 (excellent) combined_error = 0.10 → score = 80.3 (good) combined_error = 0.15 → score = 71.9 (fair boundary) combined_error = 0.20 → score = 64.4 (fair — investigate) combined_error = 0.25 → score = 57.7 (poor) combined_error = 0.30 → score = 51.7 (poor) combined_error = 0.40 → score = 41.5 (critical) combined_error = 0.50 → score = 33.3 (critical) Implication for RCA thresholds: score < 75 → combined_error > 0.13 → Warning threshold score < 60 → combined_error > 0.23 → Critical threshold
STEP 5 — CYCLE SCORE
Weighted average of three stage scores: cycle_score = S1_score × 0.20 + S2_score × 0.30 + S3_score × 0.50 Exhaust is NOT included in cycle score (pass/fail only — see Sheet 8).
RCA SUB-WINDOW DECOMPOSITION
The RCA engine does NOT re-run the full scoring formula. It decomposes each stage's normalised data into sub-windows to locate the deviation: Stage 1 sub-windows: MAE_s1_ramp = MAE during rising portion (P_norm rising from 0 → peak) MAE_s1_peak = MAE near peak (P_norm within 5% of max) MAE_s1_release = MAE during partial release (peak → local minimum) Stage 2 sub-windows: MAE_s2_ramp = MAE during second rise MAE_s2_peak = MAE near second peak MAE_s2_release = MAE during second partial release Stage 3 sub-windows: MAE_s3_ramp = MAE during third rise (S2 min → S3 peak) MAE_s3_hold = MAE during holding period (most critical) Sub-window MAEs are used as RCA rule inputs — see Sheet 3 RULES_MASTER. Sub-window boundaries are defined by dP/dt sign changes on the normalised curve.
DIVERGENCE PATTERN CLASSIFICATION
Based on which sub-window MAE is elevated, classify the deviation pattern: Pattern A — Peak too low: MAE_sX_peak > threshold, MAE_sX_ramp normal Pattern B — Slow ramp: MAE_sX_ramp > threshold (curve right-shifted) Pattern C — Oscillating: RMSE_sX >> MAE_sX (spike penalty; std_dev high) Pattern D — Release abnormal: MAE_sX_release > threshold (too shallow or deep) Pattern E — Holding drift: MAE_s3_hold > threshold (hold decays or unstable) Pattern F — Systemic: ≥2 sterilizers show same pattern in same window Pattern C detection: osc_ratio = RMSE_stage / MAE_stage If osc_ratio > 1.3 → oscillation dominant → Pattern C

## 2_STAGE_DEFINITION
Stage Definition — Triple Peak Window Boundaries (pressure-event based)
Stage_ID | Stage_Name | Start_Condition | End_Condition | Sub-windows for RCA (MAE decomposition) | Benchmark_Field (InfluxDB) | Score_Weight | Engineering_Purpose
S1 | Stage 1 First Peak | P_norm rises above 0.05 after door-closed event (dP/dt turns positive from ~0) | P_norm reaches first local minimum after first peak (dP/dt: negative → positive AND P_norm < 0.30) | MAE_s1_ramp : 0 → first peak MAE_s1_peak : within 5% of P_norm max MAE_s1_release: first peak → S1 minimum | benchmark_norm[] (full cycle, sliced to S1 window) | 20% | Air purging. Removes air pockets that reduce steam penetration. Partial release at end expels condensate and trapped air. Deviation here propagates to S2 quality.
S2 | Stage 2 Second Peak | P_norm rises again from S1 minimum (dP/dt turns positive after S1 end condition met) | P_norm reaches second local minimum after second peak (dP/dt: negative → positive AND P_norm < 0.30) | MAE_s2_ramp : S1 min → second peak MAE_s2_peak : within 5% of second P_norm max MAE_s2_release: second peak → S2 minimum | benchmark_norm[] (full cycle, sliced to S2 window) | 30% | Fruitlet loosening from FFB bunch. Heat and pressure cycling breaks bonds. Quality here directly determines stripping efficiency and oil loss. Most frequent source of low scores based on observed data (AVG 65.8).
S3 | Stage 3 Third Peak + Hold | P_norm rises from S2 minimum (dP/dt turns positive after S2 end condition met) | Exhaust valve open signal OR sustained P_norm drop > 0.05/min for > 2 consecutive minutes | MAE_s3_ramp : S2 min → third peak MAE_s3_hold : third peak plateau → exhaust start (most critical sub-window) | benchmark_norm[] (full cycle, sliced to S3 window) | 50% | Full sterilization and lipase enzyme inactivation. Holding at target pressure ensures complete heat penetration. Directly determines FFA level and oil extraction yield. Highest weight.
EX | Exhaust (Safety Only) | Exhaust valve open signal OR start of sustained P_norm drop from S3 | P_actual ≤ 0.05 bar (safe-to-open threshold) | No MAE decomposition. Pass/fail check only: - Residual pressure - Exhaust rate - Completion time | benchmark_norm[] (reference for rate check only) | 0% — not in score | Safety pressure release. Does not affect oil loss. Failure triggers safety alarm and blocks door-open command. See Sheet 8 for checks.

## 3_RULES_MASTER
Rules Master — MAE/RMSE sub-window based RCA rules (k=2.2, 50/50 weights)
Pattern: A=Peak too low ¦ B=Slow ramp ¦ C=Oscillating (RMSE>>MAE) ¦ D=Partial release abnormal ¦ E=Holding drift ¦ F=Systemic (≥2 STs same pattern) ¦ Thresholds on 0–1 normalised scale. combined_error>0.13→warn, >0.23→crit (score<75 / score<60)
Rule_ID | Pri | Stage | Attribution | Pattern | Input Signal(s) [sub-window MAE/RMSE] | Executable Formula [normalised scale 0–1] | Warn Threshold | Critical Threshold | Score Warn | Score Crit | Cap | Rec_Key → Sheet 7 | ML Label
S1-001 | P1 | S1 | Boiler | A+F | MAE_s1_peak [InfluxDB: benchmark_norm[], ster_X_norm[] sliced to S1 peak sub-window] | MAE_s1_peak = mean(¦P_norm(t) − B_norm(t)¦) for t in S1 peak sub-window FLAG if MAE_s1_peak > threshold AND ≥2 sterilizers show same pattern (Pattern F) | 0.13 < MAE ≤ 0.23 (score 64–75) | MAE > 0.23 (score < 64) | -5 | -10 | 25 | BOILER_S1_LOW_PEAK | Boiler
S1-002 | P1 | S1 | Boiler | B+F | MAE_s1_ramp [InfluxDB: benchmark_norm[], ster_X_norm[] sliced to S1 ramp sub-window] | MAE_s1_ramp = mean(¦P_norm(t) − B_norm(t)¦) for t in S1 ramp sub-window (0 → first peak) ramp_lag: actual curve right-shifted vs benchmark i.e. mean(error(t)) < 0 over ramp window FLAG if MAE_s1_ramp > threshold AND mean(error) < −0.05 AND ≥2 STs same direction | 0.13 < MAE ≤ 0.23 | MAE > 0.23 | -5 | -10 | 25 | BOILER_S1_SLOW_RAMP | Boiler
S1-003 | P1 | S1 | Boiler | C+F | MAE_s1_full, RMSE_s1_full [full S1 window] | osc_ratio = RMSE_s1_full / MAE_s1_full combined_error_s1 = 0.5×MAE_s1 + 0.5×RMSE_s1 FLAG if osc_ratio > 1.30 (RMSE >> MAE = oscillation) AND combined_error > threshold AND ≥2 STs oscillating | osc_ratio > 1.30 AND combined > 0.13 | osc_ratio > 1.50 AND combined > 0.23 | -3 | -7 | 25 | BOILER_S1_OSCILLATION | Boiler
S1-004 | P1 | S1 | Boiler | A | MAE_s1_peak, active_ster_count [InfluxDB + cycle_events] | All active STs: avg_MAE_s1_peak = mean(MAE_s1_peak across all active STs) capacity_flag = avg_MAE_s1_peak > threshold AND active_ster_count ≥ N FLAG if boiler capacity likely saturated | avg_MAE > 0.13 AND N_active ≥ 3 | avg_MAE > 0.23 AND N_active ≥ 4 | -5 | -10 | 25 | BOILER_S1_CAPACITY | Boiler
S1-005 | P2 | S1 | BPV | A+F | MAE_s1_peak [same as S1-001 but Boiler rules all PASS] | MAE_s1_peak > threshold AND Boiler rules S1-001~004 all below Warning → Boiler OK but output still low → BPV under-delivering AND ≥2 STs affected | 0.13 < MAE ≤ 0.23 | MAE > 0.23 | -5 | -10 | 20 | BPV_S1_LOW | BPV
S1-006 | P2 | S1 | BPV | C | MAE_s1_full, RMSE_s1_full [BPV pressure or downstream ster pressure] | osc_ratio_bpv = RMSE_s1 / MAE_s1 FLAG if osc_ratio_bpv > 1.30 AND Boiler rules all PASS → BPV hunting causing downstream oscillation | osc_ratio > 1.30 AND combined > 0.13 | osc_ratio > 1.50 AND combined > 0.23 | -5 | -10 | 20 | BPV_S1_HUNTING | BPV
S1-007 | P3 | S1 | Competition | A+B | MAE_s1_ramp, MAE_s1_peak [InfluxDB + ster_ramp_start_times[]] | N_conc = count(ramp_start within ±5 min window) If N_conc ≥ 2: Check if MAE_s1_ramp OR MAE_s1_peak elevated for the later-starting sterilizers (earlier ST gets steam first; later ST is starved) FLAG if N_conc ≥ 2 AND affected ST MAE > threshold | N_conc=2 AND MAE > 0.13 | N_conc≥3 OR MAE > 0.23 | -8 | -15 | 20 | COMP_S1_CONCURRENT | Competition
S1-008 | P3 | S1 | Competition | B | MAE_s1_ramp [all active STs in same window] | overlap_STs = list of STs whose S1 ramp windows overlap For overlapping STs: compare MAE_s1_ramp before vs during overlap FLAG if MAE_s1_ramp increases during overlap period | MAE increase > 0.08 during overlap | MAE increase > 0.15 during overlap | -5 | -10 | 20 | COMP_S1_OVERLAP | Competition
S1-009 | P4 | S1 | Network | A+F | MAE_s1_peak [all STs in same time window] | low_count = count(MAE_s1_peak > 0.13) in same S1 window FLAG if low_count ≥ 2 AND no concurrent ramp start event (Competition PASS) AND Boiler+BPV PASS | low_count = 2 | low_count ≥ 3 | -5 | -10 | 20 | NETWORK_S1_MULTI_LOW | Network
S1-010 | P5 | S1 | Local | A | MAE_s1_peak [this ST vs peer STs in same window] | peer_avg_MAE = mean(MAE_s1_peak of all other active STs) local_dev = MAE_s1_peak_thisUnit − peer_avg_MAE FLAG if local_dev > threshold AND only 1 ST affected | local_dev > 0.10 | local_dev > 0.18 | -5 | -10 | 15 | LOCAL_S1_UNIT | Local
S1-011 | P5 | S1 | Local | D | MAE_s1_release [S1 partial release sub-window only] | MAE_s1_release = mean(¦P_norm(t) − B_norm(t)¦) for t in S1 release sub-window (first peak → S1 minimum) FLAG if MAE_s1_release > threshold AND other sub-windows normal | 0.13 < MAE_release ≤ 0.23 | MAE_release > 0.23 | -5 | -10 | 15 | LOCAL_S1_RELEASE | Local
S2-001 | P1 | S2 | Boiler | A+F | MAE_s2_peak [S2 peak sub-window, all active STs] | MAE_s2_peak = mean(¦P_norm(t) − B_norm(t)¦) for t in S2 peak sub-window FLAG if MAE_s2_peak > threshold AND ≥2 STs same pattern | 0.13 < MAE ≤ 0.23 | MAE > 0.23 | -5 | -10 | 25 | BOILER_S2_LOW_PEAK | Boiler
S2-002 | P1 | S2 | Boiler | B+F | MAE_s2_ramp [S2 ramp: S1 minimum → second peak] | MAE_s2_ramp = mean(¦P_norm(t) − B_norm(t)¦) for t in S2 ramp sub-window mean(error(t)) < −0.05 → right-shifted (slow ramp) FLAG if MAE_s2_ramp > threshold AND mean(error) < −0.05 AND ≥2 STs same direction | 0.13 < MAE ≤ 0.23 | MAE > 0.23 | -5 | -10 | 25 | BOILER_S2_SLOW_RAMP | Boiler
S2-003 | P1 | S2 | Boiler | A | MAE_s2_peak, active_ster_count, holding_ster_count | Total demand check during S2: N_demand = N_s2_ramp + N_s3_holding avg_MAE_s2 = mean(MAE_s2_peak across all S2-active STs) FLAG if avg_MAE_s2 > threshold AND N_demand ≥ threshold | avg_MAE > 0.13 AND N_demand ≥ 4 | avg_MAE > 0.23 AND N_demand ≥ 5 | -8 | -15 | 25 | BOILER_S2_CAPACITY | Boiler
S2-004 | P2 | S2 | BPV | A+C | MAE_s2_full, RMSE_s2_full [Boiler rules S2-001~003 all PASS] | combined_s2 = 0.5×MAE_s2 + 0.5×RMSE_s2 osc_ratio_s2 = RMSE_s2 / MAE_s2 FLAG if (combined_s2 > threshold OR osc_ratio_s2 > 1.30) AND Boiler PASS AND ≥2 STs affected | combined > 0.13 OR osc_ratio > 1.30 | combined > 0.23 OR osc_ratio > 1.50 | -5 | -10 | 20 | BPV_S2_UNSTABLE | BPV
S2-005 | P3 | S2 | Competition | A+B | MAE_s2_ramp, MAE_s2_peak [+ ster_ramp_start_times[] during S2 window] | new_ramp_during_s2 = count(ramp_start while any ST in S2) If new_ramp_during_s2 ≥ 1: Check if S2-active STs show MAE_s2_ramp or MAE_s2_peak spike after the new ramp start timestamp FLAG if spike_MAE > threshold | spike_MAE > 0.10 | spike_MAE > 0.18 | -8 | -15 | 20 | COMP_S2_DEMAND | Competition
S2-006 | P4 | S2 | Network | A+F | MAE_s2_peak [all STs in same S2 window] | low_count_s2 = count(MAE_s2_peak > 0.13) in same window FLAG if low_count_s2 ≥ 2 AND Competition PASS (no new ramp event) AND Boiler+BPV PASS | low_count = 2 | low_count ≥ 3 | -5 | -10 | 20 | NETWORK_S2_MULTI_LOW | Network
S2-007 | P5 | S2 | Local | A | MAE_s2_peak [this ST vs peer STs] | peer_avg_MAE_s2 = mean(MAE_s2_peak of all other STs in same window) local_dev_s2 = MAE_s2_peak_thisUnit − peer_avg_MAE_s2 FLAG if local_dev_s2 > threshold AND only 1 ST affected | local_dev > 0.10 | local_dev > 0.18 | -5 | -10 | 15 | LOCAL_S2_UNIT | Local
S2-008 | P5 | S2 | Local | D | MAE_s2_release [S2 partial release sub-window only] | MAE_s2_release = mean(¦P_norm(t) − B_norm(t)¦) for t in S2 release sub-window FLAG if MAE_s2_release > threshold AND other S2 sub-windows normal | 0.13 < MAE_release ≤ 0.23 | MAE_release > 0.23 | -5 | -10 | 15 | LOCAL_S2_RELEASE | Local
S3-001 | P1 | S3 | Boiler | A+F | MAE_s3_hold [S3 holding sub-window — MOST CRITICAL] | MAE_s3_hold = mean(¦P_norm(t) − B_norm(t)¦) for t in S3 hold sub-window (after third peak plateaus) FLAG if MAE_s3_hold > threshold AND ≥2 STs affected | 0.13 < MAE ≤ 0.23 | MAE > 0.23 | -8 | -15 | 25 | BOILER_S3_LOW | Boiler
S3-002 | P1 | S3 | Boiler | C+F | MAE_s3_hold, RMSE_s3_hold [S3 hold sub-window] | osc_ratio_s3 = RMSE_s3_hold / MAE_s3_hold FLAG if osc_ratio_s3 > 1.30 (Boiler oscillating during hold) AND ≥2 STs show oscillation | osc_ratio > 1.30 AND combined > 0.13 | osc_ratio > 1.50 AND combined > 0.23 | -8 | -15 | 25 | BOILER_S3_HOLD_OSC | Boiler
S3-003 | P2 | S3 | BPV | E | MAE_s3_hold, RMSE_s3_hold [Boiler S3 rules PASS] | combined_s3_hold = 0.5×MAE_s3_hold + 0.5×RMSE_s3_hold osc_ratio_s3_bpv = RMSE_s3_hold / MAE_s3_hold FLAG if (combined > threshold OR osc_ratio > 1.30) AND Boiler PASS | combined > 0.13 OR osc_ratio > 1.30 | combined > 0.23 OR osc_ratio > 1.50 | -8 | -15 | 20 | BPV_S3_HOLD | BPV
S3-004 | P3 | S3 | Competition | E | MAE_s3_hold [+ ster_ramp_start_times[] during S3 window] | new_ramp_during_hold = count(ramp_start while any ST in S3 hold) If new_ramp_during_hold ≥ 1: Compute MAE_s3_hold before and after new ramp timestamp hold_MAE_increase = MAE_after − MAE_before FLAG if hold_MAE_increase > threshold | hold_MAE_increase > 0.08 | hold_MAE_increase > 0.15 | -8 | -15 | 20 | COMP_S3_RAMP_DURING_HOLD | Competition
S3-005 | P5 | S3 | Local | E | MAE_s3_hold [this ST; Boiler+BPV+Comp PASS] | MAE_s3_hold for this unit > threshold peer_avg_s3_hold = mean(MAE_s3_hold of all other STs) local_hold_dev = MAE_s3_hold_thisUnit − peer_avg_s3_hold FLAG if local_hold_dev > threshold AND only 1 ST affected | local_dev > 0.10 | local_dev > 0.18 | -8 | -15 | 15 | LOCAL_S3_HOLD_DECAY | Local
S3-006 | P5 | S3 | Local | C+E | MAE_s3_hold, RMSE_s3_hold [this ST only] | osc_ratio_local = RMSE_s3_hold / MAE_s3_hold FLAG if osc_ratio_local > 1.30 AND Boiler+BPV PASS AND only 1 ST affected | osc_ratio > 1.30 AND combined > 0.13 | osc_ratio > 1.50 AND combined > 0.23 | -5 | -10 | 15 | LOCAL_S3_HOLD_UNSTABLE | Local
S3-007 | P4 | S3 | Network | E+F | MAE_s3_hold [all STs in same hold window] | all_hold_low = count(MAE_s3_hold > 0.13) FLAG if all_hold_low ≥ 2 AND Boiler+BPV+Competition PASS | all_hold_low = 2 | all_hold_low ≥ 3 | -8 | -15 | 20 | NETWORK_S3_MULTI_DROP | Network
CY-001 | P1 | CY | Boiler | F | boiler_rule_triggers[] rolling last 10 cycles | boiler_recur = count(S1-001~004 OR S2-001~003 OR S3-001~002 triggered) / 10 FLAG if boiler_recur > threshold | recur 20–40% | recur > 40% | -8 | -15 | 30 | BOILER_CHRONIC | Boiler
CY-002 | P2 | CY | BPV | C+E | bpv_rule_triggers[] rolling last 10 cycles | bpv_recur = count(S1-005~006 OR S2-004 OR S3-003 triggered) / 10 | recur 20–40% | recur > 40% | -8 | -15 | 30 | BPV_CHRONIC | BPV
CY-003 | P3 | CY | Competition | A+B | competition_rule_triggers[] rolling last 10 cycles | comp_recur = count(S1-007~008 OR S2-005 OR S3-004 triggered) / 10 | recur 20–40% | recur > 40% | -8 | -15 | 30 | COMP_CHRONIC | Competition
CY-004 | P4 | CY | Network | F | network_rule_triggers[] rolling last 10 cycles | net_recur = count(S1-009 OR S2-006 OR S3-007 triggered) / 10 | recur 20–40% | recur > 40% | -5 | -10 | 30 | NETWORK_CHRONIC | Network
CY-005 | P5 | CY | Local | A+D+E | local_rule_triggers[] rolling last 10 cycles for this unit | local_recur = count(S1-010~011 OR S2-007~008 OR S3-005~006 triggered) / 10 | recur 20–40% | recur > 40% | -5 | -10 | 30 | LOCAL_CHRONIC | Local

## 4_THRESHOLD_LIBRARY
Threshold Library — all on normalised 0–1 scale (post normalisation)
All thresholds are on 0–1 normalised scale (after full-cycle min-max normalisation). combined_error = 0.5×MAE + 0.5×RMSE. score=100×e^(−2.2×combined): combined>0.13 → score<75 (warn), combined>0.23 → score<60 (crit).
Rule_ID | Metric_Name | Unit | Warn_Low | Warn_High (Critical starts here) | Critical_Value | Derivation / Notes
S1-001 | MAE_s1_peak | normalised 0–1 | 0.13 | 0.23 | 0.23 | score<75 at 0.13, score<60 at 0.23 (k=2.2 exponential)
S1-002 | MAE_s1_ramp + mean(error)<−0.05 | normalised 0–1 | 0.13 | 0.23 | 0.23 | right-shift confirmed by negative mean error
S1-003 | osc_ratio = RMSE/MAE | ratio | 1.30 | 1.50 | 1.50 | >1.30 = oscillation dominant vs average deviation
S1-004 | avg_MAE_s1_peak across STs | normalised 0–1 | 0.13 | 0.23 | 0.23 | N_active ≥ 3 for warn, ≥ 4 for crit
S1-005 | MAE_s1_peak (BPV context) | normalised 0–1 | 0.13 | 0.23 | 0.23 | Same as S1-001 but Boiler PASS confirmed
S1-006 | osc_ratio BPV context | ratio | 1.30 | 1.50 | 1.50 | BPV hunting → downstream oscillation
S1-007 | MAE_s1_ramp/peak + N_conc | normalised + count | MAE>0.13, N=2 | MAE>0.23, N≥3 | MAE>0.23 OR N≥3 | Competition only triggered if N_conc ≥ 2
S1-008 | MAE_s1_ramp increase during overlap | normalised delta | 0.08 | 0.15 | 0.15 | Delta MAE before vs during ramp overlap
S1-009 | low_count (MAE_s1_peak > 0.13) | integer | 2 | 3 | 3 | Count of STs with peak MAE above warn threshold
S1-010 | local_dev = thisUnit MAE − peer avg | normalised delta | 0.10 | 0.18 | 0.18 | How much worse this unit vs peers
S1-011 | MAE_s1_release | normalised 0–1 | 0.13 | 0.23 | 0.23 | Partial release sub-window only
S2-001 | MAE_s2_peak | normalised 0–1 | 0.13 | 0.23 | 0.23 | Same derivation as S1-001 for S2 window
S2-002 | MAE_s2_ramp + mean(error)<−0.05 | normalised 0–1 | 0.13 | 0.23 | 0.23 | S2 ramp from S1 minimum → second peak
S2-003 | avg_MAE_s2_peak + N_demand | normalised + count | 0.13, N≥4 | 0.23, N≥5 | 0.23, N≥5 | N_demand = N_s2_ramp + N_s3_holding
S2-004 | combined_s2 OR osc_ratio_s2 | normalised / ratio | 0.13 / 1.30 | 0.23 / 1.50 | 0.23 / 1.50 | BPV context — Boiler PASS
S2-005 | spike_MAE after new ramp start | normalised delta | 0.10 | 0.18 | 0.18 | MAE spike in S2-active STs after new ramp timestamp
S2-006 | low_count_s2 (MAE_s2_peak > 0.13) | integer | 2 | 3 | 3 | Network: multiple STs low, no competition
S2-007 | local_dev_s2 | normalised delta | 0.10 | 0.18 | 0.18 | This ST vs peer average in same S2 window
S2-008 | MAE_s2_release | normalised 0–1 | 0.13 | 0.23 | 0.23 | S2 partial release sub-window
S3-001 | MAE_s3_hold | normalised 0–1 | 0.13 | 0.23 | 0.23 | MOST CRITICAL — holding sub-window directly affects FFA
S3-002 | osc_ratio_s3_hold | ratio | 1.30 | 1.50 | 1.50 | Boiler oscillation during hold phase
S3-003 | combined_s3_hold OR osc_ratio | normalised / ratio | 0.13 / 1.30 | 0.23 / 1.50 | 0.23 / 1.50 | BPV context — Boiler PASS
S3-004 | hold_MAE_increase after new ramp | normalised delta | 0.08 | 0.15 | 0.15 | MAE_s3_hold before vs after new ramp start
S3-005 | local_hold_dev | normalised delta | 0.10 | 0.18 | 0.18 | This ST hold MAE vs peer average
S3-006 | osc_ratio_local_s3 | ratio | 1.30 | 1.50 | 1.50 | Local oscillation — Boiler+BPV PASS
S3-007 | all_hold_low count | integer | 2 | 3 | 3 | Network: multiple STs with MAE_s3_hold > 0.13
CY-001~005 | Chronic recurrence rate | fraction | 0.20 | 0.40 | 0.40 | Proportion of last 10 cycles with rule triggered

## 5_ATTRIBUTION_ENGINE
Attribution Engine — P1→P5 Priority Decision Tree
Step | Priority | Applies_To Stage | Condition_IF (MAE sub-window based) | Condition_AND (confirmation check) | Root_Cause_Output | Deduction_Cap
1 | P1 — Boiler | S1, S2, S3 | MAE_sX_peak OR MAE_sX_ramp OR MAE_s3_hold > Warning threshold (0.13) in ANY stage for rules S1-001/002/003/004 OR S2-001/002/003 OR S3-001/002 | Pattern F confirmed: ≥2 sterilizers show elevated MAE in same sub-window, same time window | ROOT CAUSE = BOILER Check: combustion, feed water, steam traps, boiler output vs demand | −25 max
2 | P2 — BPV | S1, S2, S3 | Boiler rules all PASS (MAE below 0.13 for all Boiler rules) AND S1-005/006 OR S2-004 OR S3-003 triggered (MAE or osc_ratio above threshold) | Boiler pressure within normal range BPV oscillation OR deficit confirmed osc_ratio > 1.30 OR combined > 0.13 | ROOT CAUSE = BPV Check: PID tuning, positioner, valve seat wear, air supply | −20 max
3 | P3 — Competition | S1, S2, S3 | Boiler PASS, BPV PASS AND S1-007/008 OR S2-005 OR S3-004 triggered (N_conc ≥ 2 or new ramp during hold) | MAE spike in affected STs correlates with concurrent ramp start timestamps (time-correlated deviation) | ROOT CAUSE = COMPETITION Enforce scheduling interlocks: stagger starts ≥10 min apart | −20 max
4 | P4 — Network | S1, S2, S3 | Boiler PASS, BPV PASS, Competition PASS AND S1-009 OR S2-006 OR S3-007 triggered (low_count ≥ 2, no concurrent start) | ≥2 STs show elevated MAE in same sub-window NO concurrent ramp start in same time window Deviation NOT time-correlated with any ramp event | ROOT CAUSE = NETWORK Inspect steam header, pipe resistance, steam traps, branch valves | −20 max
5 | P5 — Local | S1, S2, S3 | All P1–P4 PASS AND S1-010/011 OR S2-007/008 OR S3-005/006 triggered (local_dev > threshold for 1 ST only) | ONLY this sterilizer shows elevated MAE All peers normal in same sub-window No systemic pattern detected | ROOT CAUSE = LOCAL Inspect this unit: inlet valve, door seal, steam trap, control valve | −15 max
— | No Root Found | Any | All rules below Warning threshold (all sub-window MAEs < 0.13) | No elevated deviation detected | CLEAN CYCLE Log as good. No recommendation needed. | —

## 6_SCORE_MODEL
Score Model — weights, caps, bands, cycle formula
SCORING FORMULA (complete)
Normalisation | P_norm(t) = [P(t)−min(P_full)] / [max(P_full)−min(P_full)] ← full cycle
MAE per stage | MAE_sX = (1/n) × Σ ¦P_norm(t) − B_norm(t)¦ for t in stage window
RMSE per stage | RMSE_sX = √[ (1/n) × Σ (P_norm(t) − B_norm(t))² ]
Combined error | combined_sX = 0.5 × MAE_sX + 0.5 × RMSE_sX
Stage score | stage_score_X = 100 × e^(−2.2 × combined_sX) k = 2.2
Cycle score | cycle_score = S1_score×0.20 + S2_score×0.30 + S3_score×0.50
Exhaust | NOT in score. Pass/fail only → Sheet 8.
STAGE WEIGHTS — rationale
S1 = 20% | Air purging. Important but partial compensation possible if S2/S3 strong.
S2 = 30% | Fruitlet loosening. Directly affects stripping efficiency and oil loss.
S3 = 50% | Holding / sterilization. Most critical. Determines FFA and oil extraction yield.
Observed data | S2 AVG=65.8, S3 AVG=92.9 → Stage 2 is current weakest link. RCA priority = S2.
RCA DEDUCTION CAPS (per stage)
Boiler cap | −25 pts max per stage evaluation
BPV cap | −20 pts max per stage evaluation
Competition cap | −20 pts max per stage evaluation
Network cap | −20 pts max per stage evaluation
Local cap | −15 pts max per stage evaluation
Cycle chronic cap | −30 pts max at cycle-level chronic rules
Why caps? | Without caps, multiple rules from one group firing = unbounded negative. Cap ensures score stays 0–100 and attribution remains meaningful.
SCORE BANDS — oil loss context
≥ 90 — Excellent | combined_error < 0.05. Curve closely follows benchmark. Oil loss minimal.
75–89 — Good | combined_error 0.05–0.13. Minor deviation. Acceptable. Monitor trend.
60–74 — Fair ⚠ | combined_error 0.13–0.23. Noticeable deviation. Investigate. Moderate oil loss risk.
< 60 — Poor ✕ | combined_error > 0.23. Significant deviation. RCA required. High oil loss risk.
< 50 — Critical ✕✕ | combined_error > 0.30. Severe. Immediate inspection. Very high FFA and oil loss risk.
ATTRIBUTION OUTPUT (JSON)
Primary root cause | Group with highest total deduction after caps
Secondary root cause | Group with 2nd highest deduction if > 5 pts
Output structure | { "sterilizer":"ST-2", "cycle_id":"CY-001", "s1_score":78, "s2_score":62, "s3_score":91, "s1_combined_error":0.11, "s2_combined_error":0.22, "s3_combined_error":0.05, "s2_MAE":0.20, "s2_RMSE":0.24, "cycle_score":79.1, "root_cause":"COMPETITION", "affected_stage":"S2", "pattern":"A+B", "rec_key":"COMP_S2_DEMAND", "recommendation_EN":"...", "recommendation_BM":"..." }

## 7_RECOMMENDATION_MATRIX
Recommendation Matrix — lookup by Rec_Key → EN + BM operator message
Programmer: after Attribution Engine outputs rec_key, query this table. Display Recommendation_EN or _BM to operator. Include combined_error and stage in message.
Rec_Key | Stage | Root_Cause | Symptom (sub-window elevated) | Recommendation_EN | Recommendation_BM | Urgency | Target_Metric
BOILER_S1_LOW_PEAK | S1 | Boiler | MAE_s1_peak high ≥2 STs affected | Stage 1 first peak below benchmark on multiple sterilizers — Boiler output insufficient. Check boiler feed water supply, combustion efficiency and steam trap condition. Target: MAE_s1_peak < 0.13. | Puncak pertama Stage 1 di bawah benchmark pada beberapa sterilizer — output boiler tidak mencukupi. Semak bekalan air suap boiler, kecekapan pembakaran dan kondisi steam trap. Sasaran: MAE_s1_peak < 0.13. | High | Boiler pressure, MAE_s1_peak
BOILER_S1_SLOW_RAMP | S1 | Boiler | MAE_s1_ramp high mean(error)<−0.05 ≥2 STs right-shifted | Stage 1 ramp slower than benchmark — Boiler pressure recovery delayed after door close. Inspect boiler firing rate response and PRV condition. | Ramp Stage 1 lebih lambat daripada benchmark — pemulihan tekanan boiler lambat. Periksa kadar pembakaran boiler dan kondisi PRV. | High | MAE_s1_ramp, boiler ramp rate
BOILER_S1_OSCILLATION | S1 | Boiler | osc_ratio>1.30 RMSE>>MAE ≥2 STs oscillating | Stage 1 pressure oscillating — Boiler control loop unstable (RMSE much higher than MAE). Tune combustion PID. Check damper actuator and air/fuel ratio stability. | Tekanan Stage 1 berayun — gelung kawalan boiler tidak stabil (RMSE jauh lebih tinggi daripada MAE). Laras PID pembakaran. Semak aktuator damper. | Medium | osc_ratio = RMSE/MAE
BOILER_S1_CAPACITY | S1 | Boiler | avg_MAE_s1_peak high N_active≥3 | Boiler capacity insufficient for number of active sterilizers during Stage 1. Reduce concurrent active count or increase boiler capacity. | Kapasiti boiler tidak mencukupi untuk bilangan sterilizer aktif semasa Stage 1. Kurangkan bilangan aktif serentak atau tingkatkan kapasiti boiler. | High | N_active, avg_MAE_s1_peak
BPV_S1_LOW | S1 | BPV | MAE_s1_peak high Boiler PASS | Stage 1 BPV under-delivering — Boiler output normal but BPV not supplying adequate pressure. Recalibrate BPV setpoint. Inspect valve seat and actuator. | BPV Stage 1 tidak mencukupi — output boiler normal tetapi BPV tidak menghantar tekanan yang mencukupi. Kalibrasi semula titik tetap BPV. Periksa tempat duduk injap dan aktuator. | High | BPV outlet pressure, MAE_s1_peak
BPV_S1_HUNTING | S1 | BPV | osc_ratio>1.30 Boiler PASS | Stage 1 BPV hunting — valve oscillating causing unstable steam to sterilizers. Re-tune BPV PID. Check positioner air supply. Inspect for mechanical hysteresis. | BPV Stage 1 memburu — injap berayun menyebabkan stim tidak stabil. Laras semula PID BPV. Semak bekalan udara posisioner. | Medium | osc_ratio, BPV zero-cross rate
COMP_S1_CONCURRENT | S1 | Competition | MAE_s1_ramp/peak spike N_conc≥2 time-correlated | Stage 1 steam competition — ≥2 sterilizers started within 5 minutes. Later-starting unit starved of steam during ramp. Stagger starts by minimum 10 minutes. Implement scheduler interlocks. | Persaingan stim Stage 1 — ≥2 sterilizer bermula dalam masa 5 minit. Unit yang bermula kemudian kekurangan stim semasa ramp. Selenggara permulaan sekurang-kurangnya 10 minit. Laksanakan penguncian penjadualan. | High | N_conc, MAE_s1_ramp spike timestamp
COMP_S1_OVERLAP | S1 | Competition | MAE_s1_ramp increase during ramp overlap | Stage 1 ramp overlap — sustained steam competition between overlapping ramps. Enforce: new ramp only after previous ST completes Stage 1. | Pertindihan ramp Stage 1 — persaingan stim berterusan antara ramp yang bertindih. Kuatkuasakan: ramp baru hanya selepas ST sebelumnya selesai Stage 1. | Medium | Ramp overlap duration, MAE_s1_ramp delta
NETWORK_S1_MULTI_LOW | S1 | Network | MAE_s1_peak high ≥2 STs no competition event | Stage 1 network distribution problem — multiple STs receiving less steam than benchmark with no concurrent start. Inspect steam header for blockage or condensate trap failure. | Masalah pengagihan rangkaian Stage 1 — pelbagai ST menerima stim kurang daripada benchmark tanpa permulaan serentak. Periksa header stim untuk halangan atau kegagalan perangkap kondensasi. | Medium | MAE_s1_peak across all STs, header pressure
LOCAL_S1_UNIT | S1 | Local | MAE_s1_peak high only 1 ST peers normal | Stage 1 local issue — only this sterilizer underperforming vs benchmark. Inspect inlet steam valve for partial blockage or wear. Check local steam trap. | Masalah tempatan Stage 1 — hanya sterilizer ini berprestasi rendah berbanding benchmark. Periksa injap stim masuk untuk halangan separa atau haus. Semak perangkap stim tempatan. | Medium | local_dev MAE vs peers
LOCAL_S1_RELEASE | S1 | Local | MAE_s1_release high other sub-windows normal | Stage 1 partial release abnormal — release curve deviates from benchmark. Too shallow: air not fully purged → will reduce S2 steam penetration. Inspect exhaust valve actuator response time and condensate drain. | Pelepasan separa Stage 1 tidak normal — keluk pelepasan menyimpang daripada benchmark. Terlalu cetek: udara tidak ternyah sepenuhnya → akan mengurangkan penembusan stim S2. Periksa masa tindak balas aktuator injap ekzos. | Medium | MAE_s1_release, exhaust valve response
BOILER_S2_LOW_PEAK | S2 | Boiler | MAE_s2_peak high ≥2 STs affected | Stage 2 second peak below benchmark — Boiler output dropping during second ramp. Directly affects fruitlet loosening quality and oil loss. Check boiler load and steam demand balance during S2 window. | Puncak kedua Stage 2 di bawah benchmark — output boiler menurun semasa ramp kedua. Memberi kesan langsung kepada kualiti penanggalan buah dan kehilangan minyak. Semak beban boiler dan imbangan permintaan stim semasa tetingkap S2. | High | MAE_s2_peak, boiler output during S2
BOILER_S2_SLOW_RAMP | S2 | Boiler | MAE_s2_ramp high mean(error)<−0.05 ≥2 STs | Stage 2 ramp slower than benchmark — Boiler slow to recover after Stage 1 partial release. Check boiler response time and PRV settings. | Ramp Stage 2 lebih lambat daripada benchmark — boiler lambat pulih selepas pelepasan separa Stage 1. Semak masa tindak balas boiler dan tetapan PRV. | High | MAE_s2_ramp, boiler recovery rate
BOILER_S2_CAPACITY | S2 | Boiler | avg_MAE_s2_peak high N_demand≥4 | Stage 2 boiler capacity shortage — S2 ramps plus S3 holding units exceed boiler output. Stagger S2 starts to reduce simultaneous peak demand. | Kekurangan kapasiti boiler Stage 2 — ramp S2 ditambah unit penahanan S3 melebihi output boiler. Selenggara permulaan S2 untuk mengurangkan permintaan puncak serentak. | High | N_demand, avg_MAE_s2_peak
BPV_S2_UNSTABLE | S2 | BPV | combined_s2>0.13 OR osc_ratio>1.30 Boiler PASS | Stage 2 BPV unstable — pressure regulation poor during second peak. Impacts fruitlet loosening uniformity. Inspect BPV diaphragm and actuator wear. | BPV Stage 2 tidak stabil — pengawalan tekanan lemah semasa puncak kedua. Memberi kesan kepada keseragaman penanggalan buah. Periksa diafragma dan haus aktuator BPV. | High | combined_s2, osc_ratio_s2
COMP_S2_DEMAND | S2 | Competition | MAE_s2 spike correlated with new ramp start | Stage 2 steam demand collision — new sterilizer ramp started while another in Stage 2. Second peak quality reduced, affecting fruitlet loosening. Protect S2 window: no new ramp when any sterilizer is in Stage 2. | Pelanggaran permintaan stim Stage 2 — ramp sterilizer baru bermula semasa yang lain dalam Stage 2. Kualiti puncak kedua berkurangan. Lindungi tetingkap S2: tiada ramp baru apabila mana-mana sterilizer dalam Stage 2. | High | MAE_s2 spike timestamp vs ramp_start
NETWORK_S2_MULTI_LOW | S2 | Network | MAE_s2_peak high ≥2 STs no new ramp | Stage 2 network distribution issue — multiple STs receiving less steam in second peak without concurrent start. Inspect header pipe and steam trap condition. | Masalah pengagihan rangkaian Stage 2 — pelbagai ST menerima stim kurang dalam puncak kedua tanpa permulaan serentak. Periksa paip header dan kondisi perangkap stim. | Medium | MAE_s2_peak distribution, header condition
LOCAL_S2_UNIT | S2 | Local | MAE_s2_peak high only 1 ST peers normal | Stage 2 local issue — this sterilizer's second peak below benchmark while peers normal. Directly impacts fruitlet loosening for this vessel. Check inlet control valve and inspect for internal blockage. | Masalah tempatan Stage 2 — puncak kedua sterilizer ini di bawah benchmark. Memberi kesan langsung kepada penanggalan buah untuk vesel ini. Semak injap kawalan masuk dan periksa halangan dalaman. | High | local_dev_s2, inlet valve position
LOCAL_S2_RELEASE | S2 | Local | MAE_s2_release high other S2 sub-windows normal | Stage 2 partial release abnormal for this unit. If too shallow: fruitlets insufficiently shocked, loosening quality reduced. Inspect exhaust valve and condensate drain. | Pelepasan separa Stage 2 tidak normal untuk unit ini. Jika terlalu cetek: buah tidak tergegar dengan mencukupi, kualiti penanggalan berkurangan. Periksa injap ekzos dan saliran kondensasi. | High | MAE_s2_release vs benchmark
BOILER_S3_LOW | S3 | Boiler | MAE_s3_hold high ≥2 STs affected | Stage 3 holding below benchmark — CRITICAL. Boiler cannot maintain hold pressure. Directly reduces sterilization effectiveness, increases FFA and oil loss. Immediately check boiler output and reduce concurrent sterilizer load. | Penahanan Stage 3 di bawah benchmark — KRITIKAL. Boiler tidak dapat mengekalkan tekanan penahanan. Mengurangkan keberkesanan sterilisasi, meningkatkan FFA dan kehilangan minyak. Segera semak output boiler. | Critical | MAE_s3_hold, FFA reading
BOILER_S3_HOLD_OSC | S3 | Boiler | osc_ratio_s3>1.30 RMSE>>MAE during hold ≥2 STs | Stage 3 holding unstable — Boiler oscillating during hold phase (RMSE >> MAE). Tune boiler control loop. Check feed water pump consistency. | Penahanan Stage 3 tidak stabil — boiler berayun semasa fasa penahanan (RMSE >> MAE). Laras gelung kawalan boiler. Semak konsistensi pam air suap. | High | osc_ratio_s3, boiler hold pressure CoV
BPV_S3_HOLD | S3 | BPV | combined_s3_hold>0.13 Boiler PASS | Stage 3 BPV regulation failure during holding — CRITICAL for sterilization quality. Cannot maintain steady hold pressure. Replace BPV positioner if hunting persists. | Kegagalan pengawalan BPV Stage 3 semasa penahanan — KRITIKAL untuk kualiti sterilisasi. Tidak dapat mengekalkan tekanan penahanan yang stabil. Ganti posisioner BPV jika pemburu berterusan. | Critical | combined_s3_hold, BPV positioner
COMP_S3_RAMP_DURING_HOLD | S3 | Competition | MAE_s3_hold increases after new ramp start | Stage 3 holding disrupted by new sterilizer ramp — CRITICAL. Implement hard scheduling lock: no new ramp while any sterilizer in Stage 3 holding. | Penahanan Stage 3 terganggu oleh ramp sterilizer baru — KRITIKAL. Laksanakan kunci penjadualan keras: tiada ramp baru semasa mana-mana sterilizer dalam penahanan Stage 3. | Critical | MAE_s3_hold before/after ramp timestamp
LOCAL_S3_HOLD_DECAY | S3 | Local | MAE_s3_hold high only 1 ST peers normal | Stage 3 holding pressure decay on this unit — CRITICAL. Possible slow steam leak. Pressure test vessel: check door gaskets, safety valve, drain connections. | Kemerosotan tekanan penahanan Stage 3 pada unit ini — KRITIKAL. Kemungkinan kebocoran stim perlahan. Uji tekanan vesel: semak gasket pintu, injap keselamatan, sambungan saliran. | Critical | local_dev MAE_s3_hold, door seal condition
LOCAL_S3_HOLD_UNSTABLE | S3 | Local | osc_ratio_local>1.30 during S3 hold 1 ST only | Stage 3 holding oscillating on this unit — control valve hunting or condensate flooding. Inspect inlet control valve. Check for condensate accumulation. | Penahanan Stage 3 berayun pada unit ini — pemburu injap kawalan atau banjir kondensasi. Periksa injap kawalan masuk. Semak pengumpulan kondensasi. | High | osc_ratio_local_s3
NETWORK_S3_MULTI_DROP | S3 | Network | MAE_s3_hold high ≥2 STs Boiler+BPV PASS | Stage 3 network holding failure — header cannot sustain hold pressure for multiple units. CRITICAL: sterilization quality at risk. Inspect main header steam traps and condensate drainage. | Kegagalan penahanan rangkaian Stage 3 — header tidak dapat mengekalkan tekanan penahanan. KRITIKAL: kualiti sterilisasi terancam. Periksa perangkap stim header utama. | Critical | MAE_s3_hold across all STs
BOILER_CHRONIC | CY | Boiler | Boiler rules >20% of last 10 cycles | Chronic Boiler issue — same boiler rules triggering repeatedly. Schedule full boiler inspection: combustion, heat exchanger, steam traps, PRV. | Masalah boiler kronik — peraturan boiler yang sama dicetuskan berulang kali. Jadualkan pemeriksaan boiler menyeluruh. | Critical | Boiler rule trigger rate
BPV_CHRONIC | CY | BPV | BPV rules >20% of last 10 cycles | Chronic BPV issue — valve failing consistently. Schedule full BPV overhaul: positioner, seat, disc, recalibration. | Masalah BPV kronik — injap gagal secara konsisten. Jadualkan servis penuh BPV. | Critical | BPV rule trigger rate
COMP_CHRONIC | CY | Competition | Competition rules >20% of last 10 cycles | Chronic scheduling problem — steam competition recurring. Overhaul sterilizer scheduler with automated start interlocking. | Masalah penjadualan kronik — persaingan stim berulang. Ubah suai penjadualan sterilizer dengan penguncian automatik. | Critical | Competition trigger rate
NETWORK_CHRONIC | CY | Network | Network rules >20% of last 10 cycles | Chronic network problem — header infrastructure degrading. Full header inspection: pipe condition, steam traps, branch valves, insulation. | Masalah rangkaian kronik — infrastruktur header merosot. Pemeriksaan header penuh. | High | Network trigger rate
LOCAL_CHRONIC | CY | Local | Local rules >20% of last 10 cycles for this unit | Chronic local issue on this sterilizer — recurring mechanical problems. Full vessel inspection: all valves, door mechanism, pressure gauge calibration. | Masalah tempatan kronik pada sterilizer ini — masalah mekanikal berulang. Pemeriksaan vesel penuh: semua injap, mekanisme pintu, kalibrasi tolok tekanan. | Critical | Local trigger rate, maintenance history

## 8_EXHAUST_SAFETY_CHECK
Exhaust Safety Check — pass/fail only. Zero effect on cycle score.
Exhaust does not affect oil loss. These are SAFETY checks only. Fail = alarm + block door-open command. No score deduction.
Check_ID | Check_Name | Condition | Pass | Fail | Alarm | Action_If_Fail
EX-001 | Residual Pressure | P_actual at exhaust_end timestamp | P ≤ 0.05 bar | P > 0.05 bar | WARNING | Do not open door. Continue exhaust. Inspect exhaust valve.
EX-002 | Exhaust Completion Time | Time for P to reach ≤ 0.05 bar vs 2× expected exhaust duration | Completed within 2× expected | Not completed within 2× expected | CRITICAL | Stop operation. Inspect exhaust valve and vent line for blockage.
EX-003 | Exhaust Rate (MAE vs benchmark) | MAE of normalised exhaust curve vs normalised benchmark exhaust | MAE_exhaust < 0.30 | MAE_exhaust ≥ 0.30 | WARNING | Inspect exhaust valve actuator. Check vent line restriction.
EX-004 | Door Open Safety Gate | Hardware interlock: door-open command gated on EX-001 AND EX-002 both PASS | Both EX-001 + EX-002 PASS | Either fails | CRITICAL | BLOCK door-open command. Must be enforced in PLC/SCADA. Not a software suggestion.